### Control Table Parameters for Wrist Rehabilitation using XM540-W270-R

### Link for E-manual: https://emanual.robotis.com/docs/en/dxl/x/xm540-w270/#control-table-data-address

- ID (Address: 7)

  - Description: Identifies the motor on the communication bus.

  - Purpose: Ensures that each actuator has a unique ID for communication.

  - Default Value: 1

  - Range: 0 ~ 252

- Operating Mode (Address: 11)

  - Description: Sets the operating mode of the motor.

  - Purpose: Enables Current-based Position Control Mode.

  - Required Value: 5 (for Current-based Position Control Mode)

  - Range: 0 ~ 16

- Temperature Limit (Address: 31)

  - Description: Sets the maximum allowable temperature before the motor shuts down.

  - Purpose: Protects the motor from overheating.

  - Default Value: 80 °C

  - Range: 0 ~ 100 °C

- Current Limit (Address: 38)

  - Description: Sets the maximum allowable current to control torque.

  - Purpose: Ensures safe force application and protects the motor from overcurrent.

  - Default Value: 2047 (maximum)

  - Range: 0 ~ 2047

- Max Position Limit (Address: 48)

  - Description: Sets the maximum position limit in terms of motor pulses.

  - Purpose: Prevents overextension of the wrist.

  - Default Value: 4095

  - Range: 0 ~ 4095

- Min Position Limit (Address: 52)

  - Description: Sets the minimum position limit in terms of motor pulses.

  - Purpose: Prevents over-flexion of the wrist.

  - Default Value: 0

  - Range: 0 ~ 4095

- Shutdown (Address: 63)

  - Description: Configures conditions to shut down the motor in case of critical errors.

  - Purpose: Ensures safety by stopping the motor during overheating, overcurrent, or voltage issues.

  - Default Value: 52 (overheating and overcurrent protection)

- Torque Enable (Address: 64)

  - Description: Enables or disables the motor torque.

  - Purpose: Activates the motor for operation or disables it for manual adjustments.

  - Default Value: 0 (disabled)

  - Range:

    - 0: Torque Disabled

    - 1: Torque Enabled

- Status Return Level (Address: 68)

  - Description: Configures the return level of status packets from the motor.

  - Purpose: Determines the communication feedback behavior.

  - Default Value: 2 (return for READ and PING instructions)

  - Range:

    - 0: No return response.

    - 1: Return for PING instruction.

    - 2: Return for READ and PING instructions.

- Position P Gain (Address: 84)

  - Description: Proportional gain of the position control loop.

  - Purpose: Adjusts the responsiveness and accuracy of position control.

  - Default Value: 800

  - Range: 0 ~ 16,383

- Profile Acceleration (Address: 108)

  - Description: Sets the acceleration profile for position control.

  - Purpose: Controls the rate of change in speed for smooth motion.

  - Default Value: 0 (no profile applied)

  - Range: 0 ~ 32,767

- Profile Velocity (Address: 112)

  - Description: Sets the velocity profile for position control.

  - Purpose: Limits the maximum speed during motion for smooth operation.

  - Default Value: 0 (no profile applied)

  - Range: 0 ~ 32,767

- Goal Position (Address: 116)

  - Description: Target position for the actuator in terms of pulses.

  - Purpose: Sets the desired position the motor should reach.

  - Default Value: 0

  - Range: 0 ~ 4,095 (depends on Max and Min Position Limits)

- Present Current (Address: 126)

  - Description: Current electrical load on the motor.

  - Purpose: Monitors the torque output by measuring current.

  - Range: 0 ~ 2,047 (2.69 mA per unit)

- Present Position (Address: 132)

  - Description: Current position of the actuator in terms of pulses.

  - Purpose: Provides feedback on the motor's position.

  - Range: 0 ~ 4,095 (depends on Max and Min Position Limits)

- Present Temperature (Address: 146)

  - Description: Current internal temperature of the motor.

  - Purpose: Provides real-time thermal monitoring for safety.

  - Range: 0 ~ 100 °C